## Data
RespiraWorks Systems Requirements
Downloaded by | Ethan Chaleff
Downloaded on | May 27 2020, 11:52 am PT
Downloaded from | https://covent-19.valispace.com/specifications/groups/21/requirements
Not Reviewed or Released
Section | identifier | title | Requirement
1_General_Product | RW-SYS-016 | Reliability | The device must be able to work continuously (100% duty cycle) for a minimum of 48 hours and a target of 2 weeks without pause for maintenance, downtime, or repair.
1_General_Product | RW-SYS-057 | Power Requirements AC | The system shall be compatible with 110V/60Hz and 220V/50HZ AC mains.
1_General_Product | RW-SYS-074 | Power Requirements DC | The ventilator shall be capable of being run from a 12V power supply
1_General_Product | RW-SYS-024 | Compatible with standard/portable O2 and air supply line. | The device shall be compatible with oxygen (DISS) and supply line interfaces standards
1_General_Product | RW-SYS-063 | Hot water humidifier | The device shall be have the ability to integrate a hot water humidifier, the humidifiers that meet ISO 80601-2-74:2017 are compatible with 22mm Respiratory Tubing per ISO 5356 standard. (Note that the fitting for standard breathing tubes and humidifiers is the same)
1_General_Product | RW-SYS-076 | Max 02 Line Pressure | The system shall be compatible with O2 pressure lines up to 440Kpa
1_General_Product | RW-SYS-136 | Gas inlet markings | Gas inlet shall be clearly labelled with gas type, color, and specified pressure ranges along with direction of flow
1_General_Product | RW-SYS-135 | Shock and Vibration | The ventilator shall comply with shock and vibration tests as outlined in 80601-2-12
1_General_Product | RW-SYS-017 | Display | The device parameters and alarms shall be readable and/or audible at a distance of 1m
1_General_Product | RW-SYS-069 | Logging: Information Storage | The ventilator shall provide non-volatile data storage for at least 5000 events in an event log, including at least the following items: - Ventilation starts and stops - All confirmed Ventilation settings - All confirmed alarm settings - All occurrences and ends of alarms with all their related actions: Inhibitions, cancellations, resets, acknowledge button presses.The ventilator shall provide a date and time stamp for all event log data. The event log shall note any change to the system’s Real Time Clock by logging the Current Date/Time followed by the new Date/Time and a unique event code indicating the change.
1_General_Product | RW-SYS-054 | Human Safe Components | The system components that come into contact with the patient airway circuit shall be compatible with the following standards ISO 18562-1 First Edition 2017-03: Biocompatibility Evaluation of Breathing Gas Pathways in Healthcare Applications - Part 1: Evaluation and Testing Within a Risk Management Process ISO 18562-2 First Edition 2017-03: Biocompatibility Evaluation of Breathing Gas Pathways in Healthcare Applications - Part 2: Tests for Emissions of Particulate Matter ISO 18562-3 First Edition 2017: Biocompatibility Evaluation of Breathing Gas Pathways in Healthcare Applications - Part 3: Tests for Emissions of Volatile Organic Compounds ISO 18562-4 First Edition 2017-03: Biocompatibility Evaluation of Breathing Gas Pathways in Healthcare Applications - Part 4: Tests for Leachables in Condensate
1_General_Product | RW-SYS-137 | Altitude operation | Ventilator shall operate at altitudes below 10k ft
1_General_Product | RW-SYS-140 | Cleaning | The ventilator construction shall continue to operate while being disinfected by standard cleaning agents (e.g. bleach)
1_General_Product | HW-024 | Anti-Asphyxiation and Anti-Barotrauma | Patient needs to be protected against both asphyxiation and barotrauma in the event of a power, mechanical, or computing failure.
2_Ventilator_modes | RW-SYS-131 | Pretest Mode | The ventilator shall have a pre-test mode where the patient inhalation port and patient exhalation port are connected and the operation of the device is checked including the capability to produce the correct pressure and flow rate, the function of valves, and the function of tidal volume measurement.
2_Ventilator_modes | RW-SYS-132 | Pressure Assist Mode | The system shall support Pressure Assist Mode. In this mode a breath cycle is defined by the control parameters below. The patient can exceed the respiratory rate by initiating additional breaths, but the breath parameters do not change.Control parameters:PIP. Peak Inspiratory Pressure.PEEP. Positive End-Exhalatory Pressure.FiO2. Fraction of Inspired Oxygen.I-time. Inspiratory (rise + dwell) time.Inspiratory sensitivity.RR. Respiratory Rate.Display parameters: PIP/PEEP/RR/FiO2/I-time.Relevant Alarms: * Low tidal volume * High tidal volume * High RR
2_Ventilator_modes | RW-SYS-133 | Pressure Support Mode | The system shall support Pressure Support Mode. In this mode the breath cycle shall be defined by the control parameters below. The breath cycle shall be triggered exclusively by the patient.Control parameters:SIP: Support Inspiratory Pressure. Like PIP, but specified as a support pressure, not a target pressure.PEEP. Positive End-Exhalatory Pressure.FiO2. Fraction of Inspired Oxygen.I-time. Inspiratory (rise + dwell) time.Inspiratory sensitivityDisplay parameters: PIP/PEEP/RR/FiO2/I-time.Relevant Alarms: * Low tidal volume * High tidal volume * Low RR (apnea) *
2_Ventilator_modes | RW-SYS-134 | Pressure Control Mode | The system shall support Pressure Control Mode. This mode shall not be featured. It is effectively a subset of Pressure Assist with zero patient inspiratory effort.
2_Ventilator_modes | RW-SYS-077 | CPAP Pressure Support mode | The system shall support a HFNC operational mode. In this mode there are no breath cycles. The operator defines the following control parameters:* constant patient pressure (2-20 cmH2O, 2cm increments)constant pressure (2-20 cmH2O)constant flow rate (up to maximum possible L/min)constant FiO2 (21-100%).Needs a child requirement: Shall interface with nasal cannulas or pillows through a standard ET tube adapter interface.
3_Control_Parameters | RW-SYS-005 | Positive End-Expiratory Pressure (PEEP) control | The device shall allow the user to establish a PEEP from 0 - 20 cm H2O with increments of 2 cm H20.
3_Control_Parameters | RW-SYS-020 | Respiratory Rate control | The device shall be able to provide a Respiratory Rate between 5 - 30 BPM with increments of 1BPM. This is available in Pressure Assist Mode (RW-SYS-132)
3_Control_Parameters | RW-SYS-021 | I/E ratio control | The device shall be able to adjust the time ratio of inspiration and expiration from 1:1 to 1:4 (Inhale to exhale). Increments of 1: .1
3_Control_Parameters | RW-SYS-032 | Tidal Volume (TV) control | The device shall be able to support tidal volumes ranging from 200-800ml with increments of 10mL
3_Control_Parameters | RW-SYS-072 | Fraction of Infused Oxygen (FI02) control | The device shall be able to adjust the fraction of inspired O2 (FiO2) in supplied air from 0.21 to as close to 1 as possible with a minimum of 0.9, adjustable in increments of 0.05.
3_Control_Parameters | RW-SYS-078 | Minimum Respiratory Rate | The system shall allow the operator to set the minimum respiratory rate of the patient. The device shall initiate a breath If respiration falls below this rate. This control is available in A/C modes.
3_Control_Parameters | RW-SYS-079 | High Tidal Volume Level Alarm Trigger | This parameter will enable the operator to set a maximum inspiration tidal volume threshold between the bounds of 80 to 3000 ml in increments of 10ml
3_Control_Parameters | RW-SYS-080 | High Pip Alarm Trigger | This parameter will enable the operator to set a PIP maximum threshold. A High priority alarm will occur if the PIP level is above the HIGH PIP level for three consecutive breaths. For pressure assisted modes the high PIP shall be automatically set to PIP +5 cmH20
3_Control_Parameters | RW-SYS-081 | High Leak Level Alarm Trigger | This parameter will enable the operator to set a maximum leak threshold between the bounds of 5 to 150 lpm.
3_Control_Parameters | RW-SYS-127 | Low Tidal Volume Alarm Trigger | This parameter will allow the operator to set a minimum inspiration tidal volume threshold between the bounds of
3_Control_Parameters | RW-SYS-128 | High Respiratory Rate Alarm Trigger | This parameter will allow the operator to set a maximum respiratory rate between the bounds of 5-30 in increments of 1 BPM
3_Control_Parameters | RW-SYS-130 | Apnea Alarm Trigger | This parameter will allow the operator to set an apnea alarm trigger between the range of 1-30s with 1s incremements.
3_Control_Parameters | RW-SYS-003 | Peak Inspiratory Pressure (PIP) | The device must be capable of supplying an inspiratory pressure from 5-55 cmH2O with increments of 1 cm H2O.
3_Control_Parameters | RW-SYS-129 |  | This parameter will allow the operator to set a minimum respiratory rate between the bounds of 2-30 in increments of 1 BPM
4_Performance_Requirements | RW-SYS-083 | Patient Pressure measurement | The ventilator shall display the current patient pressure measured at the gas outlet port. The system shall display this information as a graph with duration of 30 seconds updated at a rate > 25 Hz
4_Performance_Requirements | RW-SYS-084 | Tidal Volume measurement | The system shall display the inspiratory tidal volume in range of 0mL to 3000mL with resolution of 1mL
4_Performance_Requirements | RW-SYS-085 | I:E ratio measurement | The system shall display the ratio of the inspiration time to the expiration time in the range of 10:1 to 1:200 with display resolution of 0.1 for values under 10 and 1 for values over 10.
4_Performance_Requirements | RW-SYS-086 | PEEP measurement | The system shall measure and display the PEEP in increments of 1 cmH2O
4_Performance_Requirements | RW-SYS-087 | Instantaneous Volume Measurement | The system shall display the flow rate provided to the patient as a graph with duration of 30 seconds updated at a rate &gt;25Hz and units of liters / minute.
4_Performance_Requirements | RW-SYS-042 | Assist Control Mode Delay | During patient assist control modes, the device shall provide patient pressure within 100 ms after onset of respiratory effort (Ppat < PEEP) Pressure assist mode/ Pressure support and others
4_Performance_Requirements | RW-SYS-045 | Tidal Volume per minute | The device shall be cable of delivering a minute volume of 30 liters / minute.
4_Performance_Requirements | RW-SYS-088 | General Performance Requirements | Commanded parameters not specified in the requirements shall be within 10% of the commanded value
4_Performance_Requirements | RW-SYS-089 | Inspiration Trigger Performance | Inspiration cycle shall be triggered when XX (FLOW VS PRESSURE TRIGGER)
4_Performance_Requirements | RW-SYS-092 | Peak Inspiratory Flow Rate | The device shall be capable of a peak inspiratory flow rate of at least 80L/min (best case 120 L/min)
4_Performance_Requirements | RW-SYS-093 | Breath Rate | The output breath rate shall be within +-1 breath of the commanded value
4_Performance_Requirements | RW-SYS-094 | PEEP Pressure during the exhalation Pressure Fall | The pressure overshoot during the pressure fall phase shall not exceed 0.8*PEEP-1 cmH2O
4_Performance_Requirements | RW-SYS-095 | PEEP Pressure during the Exhalation Steady Phase | Actual PEEP shall be within PEEP setting +-1 cm H2O + 10%.
4_Performance_Requirements | RW-SYS-096 | Measured Respiratory Rate | The system shall measure the respiratory rate with an accuracy of +- 1 BPM
4_Performance_Requirements | RW-SYS-082 | Respiration Rate measurement | The ventilator shall compute and display the breath rate in increments of 1 breath/min.
4_Performance_Requirements | RW-SYS-050 |  | The device shall be capable of 4 hours (1.5 hours) [WHICH IS IT?] of battery powered operation at nominal settings.
4_Performance_Requirements | RW-SYS-138 | FiO2 measurement and display | Ventilator shall display FiO2 as measured prior to patient inhale limb.
5_Alarms | RW-SYS-097 | Alarm Backup | There shall be a redundant alarm audible at 60-80 dBa at 1 meter.
5_Alarms | RW-SYS-098 | Alarm Severities | The system shall have multiple priority alarms indicating the severity of the condition. These modes shall be: High - Likely to cause PATIENT injury or death within seconds to several minutes if uncorrected Medium - Do not cause PATIENT injury or death until at least several to many minutes have elapsed Low - May cause PATIENT injury only after many minutes to hours have passed
5_Alarms | RW-SYS-099 | Alarm Priorrities | Alarms shall be shown according to their severity. The highest severity alarm shall be displayed to the user
5_Alarms | RW-SYS-100 | Alarm Visual Indicator Distance | Alarm visual indication shall be capable of being read by the user from a distance of 1m
5_Alarms | RW-SYS-101 | Alarm Visual Indication for High Severity | High severity alarms shall be indicated with a red color and a flashing frequency of 2Hz with a 50% duty cycle
5_Alarms | RW-SYS-102 | Alarm Visual Indication for Low Severity | Low severity alarms shall be indicated with a yellow color which shall be constantly on.
5_Alarms | RW-SYS-103 | Alarm Visual Indication for Medium Severity | Medium severity alarms shall be indicated with a yellow color, and a flashing frequency of 0.5Hz with a 50% duty cycle
5_Alarms | RW-SYS-104 | Alarm Auditory Signal - High Severity | High severity alarms shall have an auditory signal consisting of 4/10 pulses with each pulse lasting 200ms with a pulse frequency of 500Hz.
5_Alarms | RW-SYS-106 | Alarm Inactivation | A press on the AUDIO PAUSE / ALARM PAUSE key shall initiate an auditory alarm signal pause of all active alarms during 60 seconds (Assuming that all the active alarms can be inhibited). The alarm activation symbol shall be displayed during the inactivation state with the symbol IEC 60417-5576
5_Alarms | RW-SYS-107 | Alarm Dismissal | The system shall have a means to dismiss alarms which states have been previously set.
5_Alarms | RW-SYS-108 | Alarms - Apnea | If no breath is triggered within the specified apnea time interval, the ventilator shall respond with a medium priority alarm with the display text "APNEA ALARM"
5_Alarms | RW-SYS-109 | Alarms - Breath Rate | When the breath rate has exceeded the breath rate setting, the device shall trigger a medium level alarm with the display text "HIGH BREATH RATE"
5_Alarms | RW-SYS-110 | Alarms - Pressure Low | The system shall trigger a low pressure alarm with a high priority in the event of a low inspiratory pressure. The text shall be displayed as "LOW INSP P". This alarm shall be automatically reset upon the first breath with inspiration pressure above the threshold
5_Alarms | RW-SYS-111 | Alarms - Pressure High | The system shall trigger a high pressure alarm with high priority in the event the pressure level goes above the high inspiratory pressure threshold for at least three consecutive breath cycles. The audible indication shall be stopped upon the first inspiratory breath that goes below the threshold, but the display shall remain active until the alarm is cancelled.
5_Alarms | RW-SYS-112 | Alarms - End Tidal Volume (VTE) | The system shall initiate a medium priority HIGH VTE alarm when the VTE measured value is above the HIGH VTE for all ventilation modes except CPAP mode The system shall initiate a medium priority Low VTE alarm when the VTE measured value is above the Low VTE for all ventilation modes except CPAP mode
5_Alarms | RW-SYS-113 | Alarms - Inspiratory Tidal Volume (VTI) | The system shall initiate a medium priority HIGH VTI alarm when the VTI measured value is above the HIGH VTI for all ventilation modes The system shall initiate a medium priority Low VTE alarm when the VTE measured value is above the Low VTE for all ventilation modes
5_Alarms | RW-SYS-114 | Alarms proximal tube disconnection | The ventilator shall initiate a medium priority alarm when the proximal tube is not connected for all ventilation modes.
5_Alarms | RW-SYS-115 | Alarm Obstruction | The ventilator shall initiate a high priority alarm if no significant volume can be delivered to the patient while the target pressure is reached
5_Alarms | RW-SYS-116 | Alarms - High Leakage | A high priority alarm shall be triggered when excessive leakage is observed in the patient circuit
5_Alarms | RW-SYS-117 | Alarms - Involuntary Shutdown | The ventilator shall initiate continuous tone alarm when shut down occurs while ventilating. The alarm shall last at least 2 min.
5_Alarms | RW-SYS-118 | Alarms - Loss of AC Mains | The ventilator shall provide a medium priority alarm if an AC power source (previously available) is lost or disconnected
5_Alarms | RW-SYS-119 | Alarms - Low Battery | A ventilator operating on AC shall provide a low priority alarm when the estimated remaining operating time is 30 minutes or below.
5_Alarms | RW-SYS-120 | Alarms - Critically Low Battery | A ventilator operating on battery shall provide a high priority alarm when the estimated remaining operating time is 10 minutes or below.
5_Alarms | RW-SYS-121 | Alarms - Battery Failure | The ventilator shall provide a high priority alarm if the battery is absent or not available, or if the battery is unchargeable.
5_Alarms | RW-SYS-122 | Alarms - Overtemperature | Abnormal temperatures concerning blower (motor temperature) power management system (ambiance close to main heat sinker) and battery (cells temperature) shall be monitored and shall fall under alarm condition and safety actions if they go out of bounds
5_Alarms | RW-SYS-068 | Multiple Alarms | The device shall be allow the user to view all alarms if multiple alarms are simultaneously present.
5_Alarms | RW-SYS-139 | O2 disconnect alarm | The ventilator shall issue a high priority alarm in the event of O2 disconnection
6_Safety | RW-SYS-010 | Pressure: Inherent Protection | The device must, through inherently safe features, prevent patient pressures from exceeding 40/80(depends on SRD) cmH2O in any operating cases or failure modes.Covent SR-20:P_plateau must not exceed 30-35 cmH2OP_peak must not exceed P_plateau + 2 cmH2OOverpressure valve opens at 60 cmH2O
6_Safety | RW-SYS-012 | Filter Exhaust/Inlet | The device shall be capable of filtering intake and exhaust to below 0.3 um (HEPA) or finer in order to limit the spread of virus from contaminated patients and to prevent the contamination of the device from patients.
6_Safety | RW-SYS-123 | Audible Acoustical limit | The ventilator shall comply with the audible acoustic pressure limites as specified in 80601-2-12
6_Safety | RW-SYS-124 | Ingress Protection | Ventilator shall have ingress protection level of IP21.IP2x = No penetration into ventilator by solid objects &gt; 12.5mm (e.g. fingers)IPx1 = Dripping water (vertically falling drops) shall have no harmful effect.Ref: www.dsmt.com/resources/ip-rating-chart/
6_Safety | RW-SYS-125 | Temperature Sensing | System shall be capable of sensing the temperature of heating elements (blower motor, battery cell during operation)
6_Safety | RW-SYS-009 | Asphyxiation Prevention | The device shall prevent pressure from dropping below -4 cm H20 and limit the resistance to 3 cmH20/l/s.
6_Safety | RW-SYS-141 | Patient antiviral filters | The ventilator design shall include antiviral filters on the patient inhale and exhale limbs.