## Slide 1: Software Quality for Medical Devices

## Slide 2: 62304 Overview
IEC 62304 – International Medical device software lifecycle standard
Mandates that the developing entity have the following:
  Software quality standards
  Software risk analysis
  Defined software development process
  Software architecture and detailed design definition
  Software release configuration management
  Software testing structure
  Software maintenance plan

## Slide 3: Deliverables from 62304
Determined by software risk class
Software documentation | Class A | Class B | Class C
Software development planning | X | X | X
Software requirements analysis | X | X | X
Software architectural design |  | X | X
Software detailed design |  |  | X
Software unit implementation | X | X | X
Software unit verification |  | X | X
Software integration and integration testing |  | X | X
Software system testing |  | X | X
Software release | X | X | X

## Slide 4: Software Classification
Controller Software is definitely a Class C software
Depending on what we end up architecting UI may be class B, but probably better to approach as class C for now

## Slide 5: Software development plan
Defines the following
  Requirements Management
  Lifecycle model – Agile/waterfall or combination
  Development Environment
    Tools – Languages Toolchain, Test automation tools, IDE etc..
    Configuration Management– git, CI
    Coding standards
    Any libraries that are necessary
  Bug/Issue reporting process
  Plans for handling software not developed internally (SOUP) – QT etc..

## Slide 6: Software Requirements Documentation
Separate from system requirements
Need to be broken down into:
Types
  Non-Functional – Hardware requirements, constraints, etc..
  Functional – What the software will do
  Performance – Specific performance requirement
    (Software shall Perform XX in XX with XX accuracy)
Areas
  States of the software
  Module functional specifications
  Usability /UI specifications
  Interface Specifications – To hardware, other software etc..

## Slide 7: Software Risk Analysis
Go through hazards that can be caused by the software
  Hazard to patient
  Hazard to clinician
  Cybersecurity
Two approaches
  Top Down – Start with hazards defined in the system hazards, and say how software can contribute (Fault Tree)
  Bottom up – Analyze each module and state how this can lead to failure (DFMEA, single fault)
  Combine both into single Analysis

## Slide 8: Software Architecture
High level description of the software before implemented
Physical Description
  Controller Software runs on controller board
  UI runs on Raspberry PI
  Connected by UART
Block Diagram of software modules
  Include the SOUP (QT, RTOS etc..)
Software Interfaces
Workflow or desired states
Trace from the architecture to the requirements

## Slide 9: Software Detailed Design
Documentation stating how it will be done (can be iterative in agile approach)
Detailed documentation
  UML charts
  Functional descriptions
  Algorithm overview
Trace from the detailed design to the requirements
  Ex. Module XX implements XX requirement

## Slide 10: Verification
Unit Testing
  Plan with acceptance criteria – Document what the unit test is testing and what the expected outputs are
  Test variable init, boundary conditions, memory management as needed
Integration Testing
  Integration Test Plan – Have plan for how we will integrate the components and test together
System Verification
  Verify that all requirements are met and that all risk mitigations work

## Slide 11: Software Maintenance
Let’s decide this later
  Who we partner with and how we distribute will affect this.
  At minimum we will have to track bugs and talk about how these get triaged.

## Slide 12: How we can Implement?
TBD – We need to come to a consensus on how to manage
Group called Innolitics developed RDM to simplify process in github. Takes yaml and creates markdown and PDF documents
  https://github.com/innolitics/rdm
My Recommendations
  Keep as much documentation as we can in the github
    Coding Standards, Toolchain definitions, tools, dependencies can be markdown pages
    Architecture Can be implemented as a wiki or use RDM above
    Software Requirements should be documented separately and referenced in all git issues. Tasks that aren’t related to software requirements specifically can be tagged with #devops, #bugfix etc..
    Detailed Design – Tricky, usually defined a separate document, if we are disciplined, we could break software up into “module” folders and place documents defining each
    Hazard analysis – Manage in excel similar to the systems hazard analysis
    Verification – Will need to define the protocols and reports
      Unit Testing – Can use generator in the comments of the unit test to generate all the documentation or potentially use issues (have done this in JIRA)
As long as we have all this and follow the process we can reformat the documents later for submission.